## Slide 1: Impedance measurements notes
20/05/2025
Poppy Joshi

## Slide 2
20-05-2025
S-parameters
We can measure the ‘S (scattering)-parameters’ using the the VNA which tells us information about the reflected and transmitted components (phase and amplitude) of the EM wave. We have a 2-port network, so the s-parameters measured are:
S11 tells you how well the input port is matched to a reference impedance (typically 50 Ω) and can be easily converted to impedance using,

## Slide 3
20-05-2025
Frequency ranges
https://www.mdpi.com/2079-6374/11/11/470
Frequency range of cell impedance spectra is from 1 Hz  10 GHz, and are catagorised into 3 groups (dispersions)
α, β, γ dispersions.
α < KHz range, tells us about cell polarisation importants range for us, tricky to measure at low frequencies due to EDLs (?)
β > a few 10’s KHz, tells us about cell size, cell membrane and cytoplasm. Most commonly used
γ > GHz arises from reorientation of water molecules.
Impedence at different frequencies tells us different information:
Low frequencies (100 KHz  1 MHz) tells us cell size
A few MHz membrane capacitance
High frequency (>> a few MHz) cyctoplasm conductivity. Cells we be dehydrated so this is N/A to us

## Slide 4
20-05-2025
Keysight E5061B ENA Vector Network Analyzer
Frequency range from 5 Hz to 3 GHz
ZDUT = Impedance of the Device Under Test, 50 = 50 Ω reference impedance, S11 is the measured parameter
1-port measurement give you S11, or S22
2-port measurement gives you all 4 s-params (S11,S12,S21,S22)
P11

## Slide 5
https://assets.testequity.com/te1/Documents/pdf/keysight/N7550A-to.pdf
20-05-2025
1-port calibration

## Slide 6
21-05-2025
2-port calibration using ECal
S11 parameter for a low pass filter DC2.5 MHz
1) Connect the ECal to the VNA. Port1  Ecal port A, VNA Port 2  Ecal port B
Ground yourself when handling the Ecal unit to avoid any ESD damage!!

## Slide 7
21-05-2025
2-port calibration using ECal
2) Select frequency range of interest using the start and stop hard keys. For the low pass filter I am testing on:
Start = 5 Hz
Stop = 5 MHz
3) Press the ‘cal’ hard key to get this menu on the screen and select Ecal.
4) Select ‘2-port cal’
6) Once calibration is complete you should see a smooth function for S11, and the inverse for S21, when the Ecal is disconnected and the low pass filter is replaced.
5) Once calibration is complete you should see flat line for S11, the more negative the better.

## Slide 8
22-05-2025
Connecting via ethernet
https://www.keysight.com/us/en/lib/software-detail/computer-software/io-libraries-suite-downloads-2175637.html